## Slide 1
X-ray Cluster 1 & 2 Training
FACTS X-ray Scientist
http://research.ntu.edu.sg/facts
Version 2020.001

## Slide 2: Contents
Background on X-rays
X-ray diffraction analysis of polycrystalline materials
Diffraction geometry: BBG vs GIXRD
Hands-on training
Other matters

## Slide 3: X-ray radiation
rays are EM waves with:
  λ = 0.01 – 10 nm
  E = 0.125 – 125 keV
http://commons.wikimedia.org/wiki/File:X-ray_applications.svg
Electromagnetic radiation spectrum
X-rays have wavelength in the same order as the interatomic distances in a crystalline material
NaCl
Space group 225, Fm-3m a = 0.562 nm

## Slide 4: X-ray safety
Dangers
ionising radiation (<100nm), includes X-rays and gamma rays
Biological effects (somatic & genetic), Carcinogen
Precautions
Exposure limits (radiation worker @ 20 mSv/yr, Public 1 mSv/yr)
Protection (distance, shielding, interlock, reduce time)
Monitoring
Radiation survey meter, Area radiation monitors
Personal dosimeters (TLD, electronic dosimeters)
Background radiation (~0.1-0.3 Sv/hr)

## Slide 5: X-ray safety in FACTS
ray instruments
  IEC 61010-1:2010 compliance for newer machines
  Safety interlocks, enclosure
  L3 license issued by NEA
  Room with card controlled access
Instrument manager
  L5 license issued by NEA
  Familiar with instrument set-up, proper operation and application
  Basic instrument trouble-shooting
User
  3 ionising radiation e-learning modules
  Personal X-ray dosimeter and dosage log

## Slide 6: Generation of X-rays
Characteristic X-ray radiation
Sealed Laboratory X-ray Tube Source
Source: “Ch.2 Experimental Setups” in Powder Diffraction Theory
and Practice
Continuous or White radiation

## Slide 7: Generation of X-rays
Continuous or White radiation
Characteristic X-ray radiation
 | Anode | Energy (keV for K) | Wavelength (nm) |  |  |  | 
 |  |  | K | K1 | K2 | K | 
 | Ag | 22.11 | 0.0560868 | 0.05594075 | 0.0563789 | 0.0497069 | 
 | Mo | 17.44 | 0.0710730 | 0.0709300 | 0.0713590 | 0.0632288 | 
 | Cu | 8.04 | 0.1541838 | 0.1540562 | 0.1544390 | 0.1392218 | 
 | Co | 6.93 | 0.1790260 | 0.1788965 | 0.179285 | 0.162079 | 
 | Cr | 5.41 | 0.2291 | 0.228970 | 0.2293606 | 0.208487 | 
K  removed using metal filters or monochromator
K2  removed as well for HRXRD

## Slide 8: X-ray diffraction analysis
Crystal
periodic arrangement of atoms in space
NaCl
Space group 225, Fm-3m
a = 0.562 nm
z
y
x
y
z
x
a = 0.562 nm
(002)
(001)
(011)
(111)
(021)
Unit cell
Crystallographic planes
hkl index, d-spacing,

## Slide 9: X-ray diffraction analysis
Bragg’s law:
n = 2dsin
y
z
x
a = 0.562 nm
(002)
(001)
(011)
(111)
(021)
Crystallographic planes
hkl index, d-spacing (hkl),

## Slide 10: X-ray diffraction analysis
X-ray diffraction pattern
2-vs-Intensity plot, peaks (hkl)
2
Intensity
(hkl)
(hkl)
(hkl)
dhkl
dhkl
dhkl
2
Bragg’s law: n = 2dsin | Diffraction vector | 
Peaks appear when hkl normal and diffraction vector are aligned parallel

## Slide 11: XRD analysis: polycrystalline
Powder sample have randomly oriented crystals exposed to the incident X-rays. Some of them are perfectly aligned parallel with the diffraction vector.
X-ray
111
Det.
2 = 29.86
022
2 = 47.3
004
2 = 69.13
Silicon powder

## Slide 12: X-ray diffraction analysis
Peak Position | 
Depends on | Main info
Periodic arrangement of atoms/molecules λ of X-ray Sample preparation | Qualitative phase analysis Lattice parameters Space group Residual stress
Peak Intensity | 
Depends on | Main info
Crystal structure λ of X-ray | Quantitative phase analysis Crystal structure Preferred orientations/texture
Peak Profile (Width & Shape) | 
Depends on | Main info
Instrument Lattice distortion Microstructure | Crystallite sizes Micro-strain
Background | 
Depends on | Main info
Sample environment i.e. air, holder Amorphous phase | Amount of amorphous phase
nλ = 2dsinθ

## Slide 13: Diffraction geometry
Bragg-Brentanno: /2 scan
GIXRD: 2 scan at fixed  or 
Powder sample
Bulk sample
Thin film (with substrate)
Surface
Thin film (without substrate)
Depth profiling

## Slide 14: XRD scan modes
https://youtu.be/HrWT2M63DbU?list=PLzN0KRQO_AyrEWBW54KFP-RRNdYF2MBoR

## Slide 15: /2 : Bragg-Brentanno geometry
Powder XRD:
Bragg-Brentanno (parafocusing) geometry
Divergent beam optics:
- Including crystals with tilted hkl normal
Angular resolution can be improved with narrow
slits
hkl
Diffraction vector
X
D
2
Div. slit

Bragg-Brentanno geometry with optics
Focus length
5/March/2019
FACTS short course, Buenconsejo
15

## Slide 16: Bragg-Brentano Geometry
sample
Goniometer Circle
Focusing Circle
Geometry has to be very precise
Sample should be on the focusing circle
Randomly oriented crystals/grains (1 – 10 microns) to prevent preferred orientation
Goniometer provides the precision and accuracy required to keep the diffraction experiment in the correct geometry
X-ray optics (Incident/diffracted) conditions the beam for optimum data collection.

## Slide 17: /2: Bragg-Brentanno geometry
X
D
Goniometer circle
Focusing circle


in-focus
out-of-focus
out-of-focus
R= 320 mm (dashed curves)
= 240 mm (solid curves)
5/March/2019
FACTS short course, Buenconsejo
17

## Slide 18: Common errors in Powder XRD experiment: Sample displacement error
sample
Focusing Circle
Displacement
Peaks shift to lower angle
Peaks shift to higher angle
Goniometer Circle

## Slide 19: Common errors in Powder XRD experiment: Transparency error
sample
Focusing Circle
Goniometer Circle
Incident X-rays penetrate too
much into the sample
Peak asymmetry towards low angle
Typical for organic samples.
Can be avoided by using parallel beam and grazing incident set-up.

## Slide 20: Scan parameters
2020
20
10 20 30 40 50 60 70 80 90 100 110 120
2Theta (deg)
X-ray Cluster 1 & 2 Training
20000
10000
0
30000
40000
50000
60000
70000
80000
Intensity (a.u.)
Scan range:
Crystal structure
Information to collect
  Phase ID
  Structure determination
  Stress-Strain analysis
  Micro-structure

## Slide 21: Scan parameters: Scan time
2020
21
10
20
30
60
70
80
0
1000
2000
3000
4000
5000
6000
7000
40 50
2Theta (deg)
Intensity (a.u.)
Time/Step = 0.1s
10
20
30
60
70
80
0
10000
20000
30000
40000
50000
60000
70000
40 50
2Theta (deg)
Intensity (a.u.)
Time/Step = 1s
1s time/step
0.1s time/step
1s time/step
0.1s time/step
10
2X0 -ray C30luster 410 & 2 T5r0aining60
70
80
Normalized Intensity (a. u.)
2Theta (deg)

## Slide 22: Scan Parameters: Step size
10
20
30
60
70
80
40 50
2Theta (deg)
Intensity (a. u.)
0.02 deg 0.05deg 0.1deg
62.6
62.8
63.0
63.6
63.8
64.0
K
K
0.02 deg
0.05 deg
63.2 63.4
2Theta (deg)
Intensity (a. u.)
0.1 deg
0.02 deg
 | 30.0 | 30.1 | 30.2 | 30.3 | 30.4 | 30.5 | 30.6 | 30.7 | 30.8
2020 | X-ray Cluster 1 & 2 Training |  |  | 2Th | eta (deg | ) |  | 22 | 
0.05 deg
0.1 deg
Intensity (a. u.)

## Slide 23: Grazing Incident XRD (GIXRD)
Source: “Ch. 4 Grazing Incidence Configuration” in Thin Film Analysis by X-ray Scattering. M. Birkholz
2020 X-ray Cluster 1 & 2 Training
23
Why Grazing Incidence? Low  shallow penetration To avoid the substrate!

## Slide 24: Examples: GIXRD versus BBG scan
60
80
100
FACTS short course, Buenconsejo
0
1E13
1E12
1E11
1E10
1E9
1E8
1E7
1000000
100000
10000
1000
40
5/March/2019
20000
40000
60000
80000
Log Intensity
2 (o)
-2 scan
substrate
(222)
fcc
(311)fcc
(220)
fcc
(200)
fcc
100000
GIXRD  =1o
Intensity
63Au-35Cu-2Al thin film
fcc-phase t = 681 nm
Substrate: SiO /Si-001
2
(111)fcc
GIXRD:
Probes the thin film region only
Relative Intensity is lower
Bragg-Brentanno (/2 scan):
Additional contribution from the
substrate
Very high intensity
How much difference in the absorption and penetration depth?
24

## Slide 25
GIXRD: X-ray penetration depth
40
60
80
100
0
1E13
1E12
1E11
1E10
1E9
1E8
1E7
1000000
100000
10000
1000
40000
20000
60000
80000
100000
Log Intensity
2 (o)
-2 scan
substrate
(222)
fcc
(311)fcc
(220)
fcc
(200)
fcc
Intensity
GIXRD  =1
o
63Au-35Cu-2Al thin film
fcc-phase t = 681 nm
2
Substrate: SiO /Si-001
(111)fcc
=1o
depth =55 nm
 > 15o
depth > 900 nm
---Note-----
Porosity, columnar grains, voids, defects in the material could reduce the effective density of the film and therefore the penetration depth could be deeper.
Thin film density could be measured using XRR.
5/March/2019 FACTS short course, Buenconsejo
25

## Slide 26: Hands-on training
X-ray safety measures in FACTS
  Dosimeter usage and e-logging
  XRD room access
Sample preparation
  Shimadzu (Cluster 1)
  Bruker & Panalytical (Cluster 2)
Standard operating procedure
  Shimadzu (Cluster 1)
  Bruker & Panalytical (Cluster 2)
Post data collection analysis
  Match software for Phase ID, Quantitative analysis

## Slide 27: X-ray safety in FACTS
Access only for XRD users
Personal Dosimeter

## Slide 28: Resources

## Slide 29: X-ray cluster
Cluster 1:
  Shimadzu Powder XRD & Thin Film XRD
Cluster 2:
  Bruker D8 Advance & Panalytical Xpert Pro
Cluster 3: Training is by appointment only
  Bruker Apex II Single Crystal XRD
  Bruker D8 Discover HRXRD (multipurpose)
  Rigaku Smartlab High-flux (9kW X-ray source) XRD
  Xenocs Nanoinxder SAXS/WAXS, GISAXS/GIWAXS
  Anton Paar SAXSess SAXS/WAXS

## Slide 30: Sample Loading
Shimadzu Powder XRD
Bulk sample holder is shorter in length by about 5mm. Need to consider this as offset when loading the sample.

## Slide 31: Sample Loading
Shimadzu Thin Film XRD
Set stage height to
8.3 – 0.966 = 7.334 mm
Turn on the pump to fix the sample on the stage

## Slide 32: X-ray cluster
Cluster 1:
  Shimadzu Powder XRD & Thin Film XRD
Cluster 2:
  Bruker D8 Advance & Panalytical Xpert Pro
Cluster 3: Training is by appointment only
  Bruker Apex II Single Crystal XRD
  Bruker D8 Discover HRXRD (multipurpose)
  Rigaku Smartlab High-flux (9kW X-ray source) XRD
  Xenocs Nanoinxder SAXS/WAXS, GISAXS/GIWAXS
  Anton Paar SAXSess SAXS/WAXS

## Slide 33: Data analysis
MATCH for Phase ID, quantification
PDF 4+, ICSD database: for checking structure and Cif files
TOPAS, EVA

## Slide 34: Important things to remember after training
Only official NTU staff with NTU staff card; students and exchange students
with NTU matriculation card are allowed to use FACTS facilities.
After training, XRD users must use > 3 hours within the first 30 days of training, or their access will be automatically deactivated. To regain access, user must wait for at least the following month for retraining. There will be charges for retraining.
If there is no usage of instrument within 365 days from the last usage, access will be automatically removed by the FOM server. Retraining with charges is applicable in this case.
For booking of instruments, users are encouraged to cancel the booking early if they do not need the slots. Charges will still apply if users cancel booking within 24 hours or fail to turn up for usage thereafter.
If a user uses an instrument booked under the name by a different user, this constitutes as unauthorized usage. Penalties including access removal to ALL FACTS equipment will apply. Users are advised to request FACTS staff to change the booking to the correct user, instead of simply allowing access to someone who needs the access.
Users will bear the cost of any equipment damage and instrument access will be revoked for any misuse.

## Slide 35: X-ray Scattering Analysis @FACTS
Single Crystal XRD
Powder XRD Grazing Incident XRD
40
60
100
120
10
100
1000
10000
100000
 = 20
331
420
400
311
222
220
111
Intensity
80
2
200
 = 0.5
GIXRD scan of Au alloy film
 = grazing incident angle
0
200
400
600
800
1000
1200
1400
1600
1800
2000
+
112 114 116 118 120
2
-
112 114 116 118 120
2
112 114 116 118 120
2
= 0 = 45 = 90
Residual Stress
Triaxial residual stress analysis of Au alloy film
0
45
90
135
180
225
270
315
0.0
8.8
18
27
35
=0-90
o

Texture
(111) PF of a Cu foil
3.2
3.4
3.6
3.8
4.0
4.2
204 film
1.4 1.5 1.6 1.7 1.8 1.9 2.0 2.1 2.2
h
l
1.0E-05
1.0E-03
1.1E-01
1.1E+01
1.1E+03
1.2E+05
204 substrate
High-Resolution XRD
Reciprocal space map of an epitaxial thin film
Precession image of a diffraction pattern
Solved crystal structure:
C (black), Al (blue), Ti (grey) and F (green)

## Slide 36: X-ray Scattering Analysis @FACTS
Nano-particle size analysis
0
0.0000
0.0001
0.0002
0.0003
N(R)
5 10 15
r (nm)
Particle size distribution
2
0.0004 SiO spheres
X-Ray Reflectivity
0.0
0.5
1.0
1.5
10-1
101
103
10
5
Intenstiy
2
SiN film
Experimental Simulated
fitting results
thickness = 200.1(1.1) nm roughness = 2.05 (0.11) nm
density = 2.23 (0.02) g/cm3
SAXS GISAXS
q
z
q
xy
0.00
0.02
0.04
0.06
1
100
10
1000
10000
100000
1000000
xy
q (A-1)
in-plane scattering
q = 0.012
1
d = 52.36 nm
Nanostructured material
We also have:
We can always discuss on:
In-situ heating/cooling stage •
High-power X-ray source (Rigaku Smartlab)
XRF for composition analysis
Wide range of software for XRD data analysis
Ad-hoc in-situ experiments with supervision
Design of experiment
Data treatment and analysis